T

cobas® 8000 modular analyzer series

Short Guide Version 2.0

Table of Contents

| Table of Contents | 1 |
| --- | --- |
| System Overview - Workflow Guide | 2 |
| Starting Inspection | 2 |
| How to select calibrations and controls | 3 |
| Requesting calibrations and controls using Overview screen | 3 |
| Requesting calibrations in Calibration Status screen | 4 |
| QC after calibration | 4 |
| Requesting controls in QC Status screen | 4 |
| Start Routine | 5 |
| End of Routine | 5 |
| How to install a new application | 6 |
| Loading new application parameters | 6 |
| Test Parameter Setting | 7 |
| Test Assignment | 7 |
| Test Key Assignment | 8 |
| Configuration in Utility | 8 |
| Set up Calibrators | 9 |
| Set up Calibrators manually (for chemistry only) | 10 |
| Set up Roche Controls | 10 |
| Edit target values and ranges for Roche controls | 11 |
| Edit target values and ranges for non Roche controls | 11 |
| Set up non Roche Controls | 12 |
| Rack Assignment for non barcoded calibrators and controls | 13 |

System Overview - Workflow Guide

Starting Inspection

Check system alarms.

Select beginning of the day maintenance. Select Clear or Backup and Clear.

Print load list.

Activate the Preventive Action check box, if not selected.

Load reagents and auxillary reagents onto the analyzer according to reagent load list. Remove all empty packs from the system.

Perform Prime for ISE. Select all.	 After the Reagent Prime for ISE, recalibrate the ISE unit before you resume routine operation.

How to select calibrations and controls

Requesting calibrations and controls using Overview screen

Calibration area

QC area

Print load list.

Load calibrator and QC racks on the analyzer according to calibrator and QC load list.

Requesting calibrations in Calibration Status screen

Select test.	Select calibration method e.g.	.	 Calibration request is active when bar is blue  .

Print load list.

QC after calibration

- Insert the QC rack directly after the calibrator racks.

Requesting controls in QC Status screen

Select test

or select    for all active Reagent Packs.

Select    for Stand by QC.	Select test

Print load list.

Start Routine

Load racks with samples.

End of Routine

Deselect rack reception mode, if necessary.

Area Rack Reception Mode

Deselect    Rack Reception Mode.

Load green rack:

Pos. 1	MultiClean Pos. 2	SysClean Pos. 3	Activator

How to install a new application

Loading new application parameters

Select one of the available search conditions.	 Activate the check box of parameter to be downloaded.

Select the unit, if necessary:                           Check if the parameters are downloaded.

c 701 / c 502: For first installation parameters for dilution and cleaner reagent packs have to be downloaded.

Test Parameter Setting

Test Assignment

Select

Optional settings    Define decimal places, select Automatic

Rerun, define Control Interval time,

AutoQC on board stability, Repeat Limit, Qualitative fields.

Select ISE module.	Select Na, K, Cl or Na, K.

Select c 701 module.	Select checkbox  ,               Select c 502 module.	Select checkbox

Select c 602 module.	Select checkbox

Test Key Assignment

Select

Select a test.

Configuration in Utility

Select option for Routine/Stat, Calibrator and Control.	 
Select checkmark and Hours.

Define cycles for wash rack: photometric: 5, ISE: 15.

Select Expired Reagent Flag.

Select Auto Masking.	Select Reagent Information Printout.

Select Print Dilution Ratio.

Set up Calibrators

Load reagent packs on the module.

Select one of the available search conditions.

Activate the check box of the parameter to be downloaded.

Set up Calibrators manually (for chemistry only)

Select an empty line.

Type in calibrator name, code, lot no., and expiration date.	    Type in the concentration values.

Set up Roche Controls

Load reagent packs on the module.

Select one of the available search conditions.	 Activate the check box of the parameter to be downloaded.

Edit target values and ranges for Roche controls

Log on to data manager. Select Control Material.

Select                  Highlight the control material you want to edit.

Select                Highlight the test you want to edit. Select                    Select .

Type in the value.

Edit target values and ranges for non Roche controls

Log on to data manager. Select Control Material.

Select              Highlight the control material you want to edit.

Select            Highlight the test you want to edit. Select	Type in the value.

Set up non Roche Controls

Log on to data manager. Select Control Material.

Select       and right-click in the work area.

Select   on the shortcut menu.

Complete the fields according to the information from the control material manufacturer.

- Control name

- Control Code (enter a number between 25000 - 32000)

- Sample type

- Lot No.

- Expiry date

- Instrument

- Test

to save the test without closing the window, or       to save the test and close the window.

Rack Assignment for non barcoded calibrators and controls

To remove the assignment

To remove the assignment

Do not mix barcoded and non-barcoded calibrators/ controls on the same rack.